## Table 1 & 2
Using a sweep generator and spectrum analyzer or signal generator and power meter
Table 1 - Detailed Crystal Measurement Procedure Using a Power Meter or Spectrum Analyzer
Step 1: Remove crystal and adjust neutralizing capacitor for minimum signal into power meter
or spectrum analyzer at the crystal frequency.
Step 2: Short out crystal and measure signal level with power meter or spectrum analyzer (P1 dBm).
Step 3: Insert crystal and measure signal level with power meter or spectrum analyzer (P2 dBm).
Step 4: Determine power difference between P1 and P2.
| Power Difference | 2.12 | dB | P1 - P2
| Rs | 10.56 | ohms | Rs = (38.2*10^([P1-P2]/20)-38.2)
Step 5: Measure crystal series resonance frequency on the power meter or spectrum analyzer.
| Fs | 5.998 | MHz
Step 6: Measure 3 dB bandwidth on power meter or spectrum analyzer.
| BW3dB Measured | 249 | Hz | Q, hence BW, is lowered by test fixture
| Q (loaded) | 24088.5221 |  | Qu = Fs*10^6/BW3dB, ref 2.
Loaded Q, includes effects of test fixture.
| Q(unloaded) | 111227.1414 |  | Take out effects of test fixture. Q = 2*PI*Fs*Lm/Rs, ref 3.
| BW3db Unloaded | 53.9261 |  | BWunloaded = Fs/Qunloaded, ref 2.
| Cm | 0.022591 | pF | Cm = BW3dB/((2*PI*(Rs + 38.2)*Fs^2), ref 2.
| Lm | 31.1662 | mH | Lm = 1/((2*PI*Fs)^2 * Cm), ref 2.
Step 7: Measure the parallel resonant frequency Fp on the spectrum analyzer.
| Fp | 6.0128 | MHz
| Co | 4.6021 | pF | Co = ((10^12)*(Cm*10^-12)*Fp*10^6)/(2*(Fp-Fs)*10^6)
Table 2 - Detailed Crystal Measurement Procedure Using a Vector Voltmeter
Step 1 | Freq set for ¦Z¦ = min and phase = 0 degrees. The ¦Z¦ reading is then Rs and the frequency is Fs.
| Rs | 12.5 | ohms
| Fs | 5.9981 | MHz
Step 2 | The frequency is changed slightly until phase = +45 degrees and F45 noted.
| F45 | 5.9981 | MHz
Step 3 | The frequency is then adjusted until phase = -45 degrees and F-45 noted.
| F-45 | 5.998 | MHz
| BW unloaded | 59 | Hz
Step 4 | Calculate Q = Fs/(F45 - F-45)
| Q (unloaded) | 101662.2373 |  | Qu = Fs*10^6/((F45-F-45)*10^6), ref 2.
| Cm = | 0.02088 | pF | Cm = 1/(2*PI()*Fs*Qu*(Rs)), ref 2.
| Lm = | 33.7193 | mH | Lm = 1/(((2*PI*Fs)^2)*Cm), ref 2.
Step 5 | Co is obtained by shifting the frequency <= 0.99Fs, reading ¦Z¦, and then computing Co.
| Z | 4600 | ohms
| F (<=.99Fs) | 5.9 | MHz
| Cr (residual cap.) | 1.4 | pF | Residual capacitance of component fixture
| Co = | 4.4642 | pF | Co = 1/(2*PI*F*Z) - Cr
References:
1. "Design of Crystal and Other Harmonic Oscillators", Benjamin Parzen with Arthur Ballato, 1983, John Wiley and Sons, Inc., Page 424 - 425
2. "Crystal Device Measurements Using the Spectrum Analyzer", Tektronix AX-3325, 2/77
3. "Crystal Filters Design, Manufacture, and Application", pg 14. Robert G. Kinsman